C4G blis EQUIPMeNT interface manual

VERSION 1.0

# PREFACE

## PURPOSE OF THIS DOCUMENT

The goal of this document is to spell out all the necessary configurations to be made on BLIS and on any equipment that needs to be interfaced (either uni-directional or bi-directional) with BLIS. This document should be consulted at all times if there are any doubts in configurations on client sites.  All supported equipment are listed in this document.  Please note that only equipment listed here can be interfaced using the java based BLIS Client Interface Tool.

## WHO SHOULD READ THIS DOCUMENT

This manual is written for BLIS (Basic Laboratory Information System) site implementers, developers and for those who want to know more about how BLIS is interfaced with equipment and other external systems in the medical facility.

For successful BLIS interface configuration and deployment at client sites, it is recommended that implementers are equipped with below skills;

| Skill | Level of Competence |
| --- | --- |
| SQL | Intermediate |
| XML | Basic |

Implementers also need to pay particular attention to details.

## WHAT CAN YOU FIND IN THIS DOCUMENT

This configuration manual employs the step by step approach in helping you identify configuration sections, you can also find in this document pictures, sample SQL queries, explanations on configuration values and sample settings which can help you to fast track your site BLIS interface implementations. You will also find vivid clarifications on configurations for all supported equipment.

## CONVENTIONS USED IN THIS DOCUMENT

The chart below explains the icons used in this manual

| When you see… | Then … |
| --- | --- |
|  | The statement following the icon is a tip or a reminder of important information you are assumed to know. You can skip this information. |
|  | The statement following the icon is important information and should not be ignored. It also alerts you to potential problems that you may encounter along the way. Read and remember these tidbits to avoid possible trouble. |
|  | This icon signifies that the accompanying explanation may be informative but isn’t crucial to BLIS interface deployment. Feel free to skip these snippets, if you like. |

## REVISION HISTORY

| Version | Date | Author | Change |
| --- | --- | --- | --- |
| 1.0 | 13/06/2014 - | Stephen Adjei-Kyei | Initial Draft |

Table of Contents

PREFACE	i

PURPOSE OF THIS DOCUMENT	i

WHO SHOULD READ THIS DOCUMENT	i

WHAT CAN YOU FIND IN THIS DOCUMENT	i

CONVENTIONS USED IN THIS DOCUMENT	ii

REVISION HISTORY	ii

THE BLIS INTERFACE CLIENT	1

HOW IT WORKS	1

Diagram 1	2

Typical bi-directional flow	2

Diagram 2	3

Typical uni-directional flow	3

Diagram 3	3

IMPORTANT CONFIGURATION FILES	4

Main Configuration file	4

Sections	5

Analyzer specific configuration files	8

Sections	8

SETTING UP BLIS INTERFACE CLIENT ON WINDOWS	11

SETTING UP BLIS INTERFACE CLIENT ON LINUX	13

SETTING UP BLIS INTERFACE CLIENT ON MAC OS	13

SPECIFIC CONFIGURATIONS FOR INTERFACING ANALYZERS	13

Mindray BS-200E	15

ABX Pentra 60 C+	17

ABX MACROS 60	20

SYSMEX XS 500i	21

BT 3000 PLUS	21

BD FACSCalibur	21

ABX Pentra 80	21

SELECTRA JUNIOR	21

Database Section	21

How to know the Working Database	22

Tables to consider	24

Structure of Equipment Interface tables	25

Database Search scenarios	27

# THE BLIS INTERFACE CLIENT

BLIS Interface client is a very lightweight tool for Interfacing BLIS with analyzers used in the medical laboratory. There are 2 main deployment modes. (Uni-directional and Bi-directional) The deployment mode actually differ per analyzer and this is purely based on the capability of the analyzer that has to be interfaced.

|  | Bi-directional Communication: 2 Way communication; Analyzer is able to send and receive information from BLIS |
| --- | --- |
|  | Uni-directional Communication: 1 way communication; Analyzer only send results to BLIS |

After a successful deployment;

- Bi-directional:

  - Patient demographics, specimen and test details on BLIS can be downloaded automatically into the analyzer

  - Test results can be transferred from Analyzer to BLIS automatically

  - There is no need to register sample and patient details on analyzer; everything is done on BLIS and transferred to analyzer.

- Uni-directional:

  - All tests performed on Analyzer with correct BLIS Specimen ID are automatically sent to BLIS.

  - BLIS demographics cannot be transferred to analyzer

|  | BLIS Client Interface tool is a JAVA based application. Communication Protocols used are RS232, TCP/IP. Also BLIS Interface Client can directly connect to analyzer database or a flat file if accessible. The Messaging formats supported are HL7, ASTM, ABX Proprietary Format, Vital Proprietary format, CLSI LIS2-A, CLSI LIS2-A2 The protocol and messaging format used is purely determined by the capability of the analyzer that has to be interfaced. |
| --- | --- |

## HOW IT WORKS

BLIS Interface Client connects to BLIS via BLIS HTTP API and pulls details of patient, specimen and tests. It convert the details into analyzer format then transmit to the analyzer using analyzer supported protocol and message format. It is a platform independent tool which can run on Windows, Linux, UNIX and Mac.

Find Flow Diagram below for pictorial view

### Diagram 1

BLIS INTERFACE FLOW DIAGRAM

### Typical bi-directional flow

A bi-directional interface with an analyzer has below steps:

- Specimen Registration on BLIS

- Specimen details are fetched from BLIS by BLIS Interface Client

- Formatted details are sent to analyzer for tests to be performed

- Tests results are sent from Analyzer to BLIS Interface Client

- BLIS Interface Client transmits results to BLIS to be saved.

### Diagram 2

### Typical uni-directional flow

A Uni-directional interface with an analyzer has below steps:

- Specimen Registration on BLIS

- Lab Technician register specimen with BLIS Specimen ID on Analyzer

- Tests results are sent from Analyzer to BLIS Interface Client

- BLIS Interface Client transmits results to BLIS to be saved.

### Diagram 3

## IMPORTANT CONFIGURATION FILES

|  | Please pay particular attention to details in these files. The behavior of BLIS Interface Client depends on them. If you have any doubts about a change please refer to the section which explains that section or contact GHSS BLIS Team for more clarification. |
| --- | --- |

### Main Configuration file

The main configuration file is BLISInterfaceClient.ini (Text Based configuration) located at the main directory of BLIS Interface Client. The format of BLISInterfaceClient.ini is like most popular configuration files for php (php.ini), mysql (my.ini), etc.

|  | Beginning a line with Semi colon (;) sets that line as a comment and will not be used by BLIS Interface Client. Example: ;****************************************************************** ; I am a comment. The above line is a comment. The line below is also a comment ;****************************************************************** |
| --- | --- |

### Sections

#### [FEED SOURCE]

#### This section specifies the source from which information will be retrieved. Supported options are

#### RS232: Connection via serial port. This option is used for equipment where TCP/IP connection is not available.

#### TCP/IP: Connection via TCP sockets with server IP address and port.

#### MSACCESS: Direct database access

#### HTTP: Connection via HTTP protocol

- TEXT: Connection to flat files for data

#### [MSACCESS CONFIGURATIONS]

This portion is used to configure settings for Microsoft Access connection.  Configuration items are:

- DATASOURCE: The name of the ODBC data source configured at the ODBC Data source administrator dialog. As of this writing, Only ABX Pentra 60 C+ supports MS Access configurations. Please refer to ABX Pentra 60 C+ configuration section for details on how to configure system data source at the ODBC Administrator dialog from control panel.

- DAYS: No of days old data to pull from BLIS. Default is 0 which means current day.

#### [RS232 CONFIGURATIONS]

This portion is used to configure settings for serial port connection. Items are:

- COMPORT: The port on which to communicate with the analyzer.  If the analyzer you want to interface uses RS232 communication, refer to the section on how to find the comport on your system.

- BAUD_RATE: The speed of connection.eg (4800,9600,38400,57600,115200)

- PARITY: How is data checked for errors? Available options are (Odd, Even, None). Default is “None”.

- STOP_BITS: Choose the time between character transmissions for your serial device. Available options are (1, 2). Default is 1

- DATA_BITS: Choose the number of data bits used by your serial device to represent characters. Available options are (7, 8). Default is 8

- APPEND_NEWLINE: Should the data transfer be appended with the Newline (/n) character (Yes/No). Default is No

- APPEND_CARRIAGE_RETURN: Should the data transfer be appended with the Enter (/r) character (Yes/No). Default is No

#### [TCP/IP CONFIGURATIONS]

This portion is used for the configuration of TCP/IP connection. Items are:

- PORT: Port on which to listen for data

- EQUIPMENT_IP: The IP address of the  analyzer system if applicable

- MODE: How should the BLIS Interface Client connect to analyzer? Should it connect as a server or client? Options (server/client). Default is server

- CLIENT_RECONNECT: This setting is used only if MODE=client. Should the application reset client connection when server sends "End of Transmission (EOT)"? Options (yes/no). Default is yes

#### [TEXT]

This portion is used to configure reading of text files for test results. Items are:

- BASE_DIRECTORY: The parent directory where all files will be located

- USE_SUB_DIRECTORIES: Should sub directories be opened for files? (Options are yes/no)

- SUB_DIRECTORY_FORMAT: Sub directories naming format. Example ddMMyy. Only used when USE_SUB_DIRECTORIES = yes

- FILE_NAME_FORMAT: Files naming format

- FILE_EXTENSION: File extension used. Example txt,csv,exp,evl

- FILE_SEPERATOR: File value separator used. Supported Options (TAB,COMMA,COLON,SEMI-COLON,SPACE)

#### [BLIS CONFIGURATIONS]

This portion is used to set the BLIS HTTP server configurations that will be used by this tool to communicate with BLIS. Items are:

- BLIS_URL: The URL on which BLIS is accessed from the web browser. Example is http://192.168.56.1:4001/

- BLIS_USERNAME: BLIS account username

- BLIS_PASSWORD: BLIS account password

#### [EQUIPMENT]

This section contains the list of all supported equipment this version of BLIS Interface Client can communicate with. This portion is pre-filled with all supported analyzers.

|  | Please note that if your equipment is not listed here you cannot connect with this version of BLIS Interface Client. To use any of the equipment listed here just uncomment that line. It is extremely important that you uncomment only ONE (1) equipment at a time. |
| --- | --- |

#### [MISCELLANEOUS]

This portion is used to set miscellaneous application configurations. Items are:

- ENABLE_LOG: Set whether you want logs to be generated by application. Options are (yes/no). Default is yes

- WRITE_TO_FILE: Set to enable or disable writing messages received from analyzer to file. Options are (yes/no). Default is no

- POOL_DAY: How many days old data to pull from BLIS. 0: is for current day, 1: for yesterday, 2: for 2 days ago, etc. Default is 0

- POOL_INTERVAL: Time interval in seconds to wait and fetch for data in BLIS. Default is 30 seconds

- ENABLE_AUTO_POOL: Set whether automatic data retrieval is enabled or not. Options are (yes/no). Default is  yes

### Analyzer specific configuration files

Every analyzer has its own xml configuration file. All specific settings are made in their respective files.

They are located at mainfolder\configs\AnalyzerName\*

The purpose of the analyzer specific configuration file is to match BLIS specimen types, test types and measures to that of the analyzer so that tests can be sent seamlessly from the analyzer to BLIS and test can be saved into BLIS from the analyzer as though they were entered manually by technicians.

Please refer to the configuration section of the particular analyzer you want to interface with BLIS for specific details.

### Sections

#### supportedtestTypes

This portion is used to configure BLIS tests that are supported by the analyzer. BLIS Interface Client will only send BLIS tests configured here to the analyzer. The most important field to fill here is the testtypeid

test_type_id from BLIS Database

You just have to repeat the testtype block for each BLIS test type you want to configure.

<texttypeid>BLIS TesttypeID</testtypeid>

You must select the BLIS test_type_ids from the database using SQL.

|  | Find example SQL query below to check and identify the test_type_id for Potassium. Query below might yield more than 1 row; look through and pick the exact test_type you want. SELECT test_type_id, name FROM `test_type` WHERE name LIKE '%Potassium%' Please note that above query is just an example. Depending on the implementation you are performing, you might write other queries to find the BLIS test_type_id. Please refer to the database section for more details |
| --- | --- |

<testtypename>BLIS TestType name</testtypename>

Type in the name of the test for which you have setup the ID.

|  | Please note that tests not configured here will not be sent to analyzer. All supported tests must be setup here. |
| --- | --- |

#### Supportedsamples

This portion is used to configure BLIS specimen types that are supported by the analyzer. BLIS Interface Client will only send BLIS specimens configured here to the analyzer. The most important element to fill here is the lissampleid

Specimen_type_id from BLIS Database

You just have to repeat the sample block for each BLIS specimen type you want to configure.

< lissampleid >BLIS specimen_type_id </ lissampleid >

You must select the BLIS specimen_type_ids from the database using SQL.

|  | Find example SQL query below to check and identify the specimen_type_id for Urine. SELECT specimen_type_id, name FROM `specimen_type` WHERE name LIKE '%Urine%' Please note that above query is just an example. Depending on the implementation you are performing, you might write other queries to find the BLIS specimen_type_id. Please refer to the database section for more details |
| --- | --- |

|  | Please note that specimen types not configured here will not be sent to analyzer. All supported specimen must be setup here. |
| --- | --- |

#### supportedtests

This portion is used to configure BLIS measures that are supported by the analyzer. BLIS Interface Client will only send BLIS measures configured here to the analyzer. The most important element to fill here is the listestid

This ID is dependent on the analyzer and is configured by developer. Please do not modify any equipmenttesttypeid. Leave it as it is.

measure_id from BLIS Database

|  | Please note that Measures in BLIS are the actual tests that are performed on the analyzer and not the Test Types. For Instance a BLIS test type like FBC could have measures like WBC, RBC, HGB, etc. These measures are what you need to configure here. |
| --- | --- |

You just have to repeat the  listest block for each BLIS measure you want to configure.

< listestid>BLIS measure_id </ listestid>

You must select the BLIS measure _ids from the database using SQL.

|  | Find example SQL query below to check and identify the measure _id for Potassium. SELECT measure_id, name FROM `measure` WHERE name LIKE '%Potassium%' Please note that above query is just an example. Depending on the implementation you are performing, you might write other queries to find the BLIS measure_id. Please refer to the database section for more details |
| --- | --- |
|  | Every measure is related to a test_type. Find example SQL query below to find all measures that belong to Electrolyte test type: SELECT tm.test_type_id, tm.measure_id, t.name AS testname, m.name AS measurename FROM `test_type_measure` tm INNER JOIN test_type t ON t.test_type_id = tm.test_type_id INNER JOIN measure m ON m.measure_id = tm.measure_id WHERE t.name LIKE 'Electrolyte' Example output of query above You need to configure all 3 measures here plus all other measures from all supported BLIS test_types. Please refer to the database section for more details |

|  | Please note that measures not configured here will not be sent to analyzer. Tests performed on analyzer will not be sent to BLIS if they have not been configured here. |
| --- | --- |

## SETTING UP BLIS INTERFACE CLIENT ON WINDOWS

- Copy the BLIS Interface Client Folder to Drive C:\

- The tool needs Java runtime to run. The main folder contains Java runtime environment 7 update 51. Install this file first.

Install first

You can get the current version of java runtime environment from here: http://java.com/en/download/

- Make configurations in the BLISInterfaceClient.ini file based on the equipment you want to interface. Please refer to the section which explains required and correct configuration for your analyzer.

- Make supported BLIS specimen types, tests types and measure configurations in the appropriate xml configuration file for your analyzer. Please refer to the section which explains required and correct configuration for your analyzer. For example Mindray BS-200E configuration can be found at mainfolder\configs\hl7\MindrayInterface.xml

- Click on Install.bat  to setup BLIS Interface Client to start when Windows starts

Just Click!

This makes the application to start when you boot up the computer

- Click on BLISInterfaceClient.jar to start tool

## SETTING UP BLIS INTERFACE CLIENT ON LINUX

## SETTING UP BLIS INTERFACE CLIENT ON MAC OS

## SPECIFIC CONFIGURATIONS FOR INTERFACING ANALYZERS

1. Put correct BLIS URL here

2. Create a new user account for the analyzer and put details here.

Below sections are required regardless of the type of analyzer used. These section must be set correctly. Find example settings below:

##### [BLIS CONFIGURATIONS]

##### [EQUIPMENT]

All supported equipment has already been listed and commented here:

Just uncomment your analyzer name

##### [MISCELLANEOUS]

Make changes as required

The various specific equipment configuration xml files are very similar. They have the same sections and so you do no need any special training in setting them up.

All the files are located at Mainfolder/configs/;   as at the time of this writing, there are 10 configurations ready for deployment.

This section will explain the specific configurations that needs to be made in the main application configuration file (BLISInterfaceClient.ini).

### Mindray BS-200E

The Mindray BS -200E is a chemistry analyzer. This analyzer uses TCP/IP connection.

|  | Please note that: The computer attached to the analyzer needs to be configured to reach the BLIS network. Depending on the facility you want to implement this; you might require a network cable or a wireless adapter. |
| --- | --- |

#### BLISInterfaceClient.ini

Configuration sections that needs to be modified are as follows:

##### [FEED SOURCE]

Set it as it is here

##### [TCP/IP CONFIGURATIONS]

Correct MODE for this analyzer is server

#### MindrayInterface.xml

This is the main configuration file for BS-200E and can be located at

\BLISInterfaceClient\configs\hl7\ MindrayInterface.xml

|  | The <Message></Message> Section should be ignored. These configurations are used by the system and was configured by developer. |
| --- | --- |

The rest of the sections are same as other xml files.

|  | The Mindray BS-200E analyzer allows you to configure all the BLIS measure IDs in the analyzer. This is to map the analyzer test to BLIS ID so that the analyzer will send the BLIS ID to BLISInterfaceClient. Please ask for more details from the LAB Technicians at the facility |
| --- | --- |

### ABX Pentra 60 C+

The ABX Pentra 60 C+ is a hematology analyzer. The BLISInterfaceClient connects to its MS Access Database.

Below are the configuration that needs to be setup in BLISInterfaceClient.ini

##### FEED SOURCE]

Set it as it is here

Data source name

#### How to setup the data source name

The ABX Pentra 60 C+ is setup on Windows XP

- Open Control Panel

- Click on Administrative tools

- Click on ODBC Data sources

- Select the System DNS Tab

Add new Data source

- Click on Add

- Select “Driver do Microsoft Access (*.mdb)”

Select MS Access driver

- Click on finish

- Enter a name for your data souce

- Click on select button to select the MS access database

Enter Data source name

Select the database

- Select the database file

  - By default the database name is ASP01DB.MDB

  - It is usually located at C:\program files\abx\data\

  - Select the file and click on OK

- Enter the data source name you created at the data source section in the configuration file as indicated above.

### ABX MACROS 60

The ABX Macros 60 analyzer has no computer attached to it. BLISInterfaceClient connects to it using RS232 protocol.

Below are the configuration that needs to be setup in BLISInterfaceClient.ini

### SYSMEX XS 500i

### BT 3000 PLUS

### BD FACSCalibur

### ABX Pentra 80

### SELECTRA JUNIOR

## Database Section

BLIS uses MySQL as its backend database. You can connect to the database using any MySQL client; like MySQL workbench, MySQL Query Browser, MySQL front or phpMyAdmin.

phpMyAdmin has been included in the BLIS package. You can connect to the database after you have started BLIS.exe. Use example URL as follows:

http://192.168.56.1:4001/phpMyAdmin

Username: root

Password: blis123

### How to know the Working Database

You need to know the current working database for the site you are interfacing before you can select accurate data. To get the site database follow steps below:

- Log into BLIS as administrator

- Click on Lab Configuration

- Look at the Address bar for the database name

Lab Config ID is 1007. Database will be blis_1007

- Open phpmyAdmin using previous section details.

- Look at the database list on the left hand side and select correct database

1

2

Select database blis_1007

3

### Tables to consider

There are 4 main tables to consider for BLIS equipment interfacing:

- specimen_type : Contains the details of all configured BLIS specimen types

- test_type : Contains the details of all configured BLIS tests types

- measure : Contains the details of all measures in BLIS

- test_type_measure: Contains the association of all BLIS measures to their respective test types

### Structure of Equipment Interface tables

### Database Search scenarios

As a BLIS equipment interface site implementer you will have to write SQL queries in the database to search for IDs to configure your equipment. Here are some example scenarios you may encounter.

|  | Please note that examples below are just for demonstration purpose. You will need to write your own queries and more complex queried as desired based on your need. |
| --- | --- |

Any word beginning with whole

Specimen Types

- Get the ID of Whole Blood specimen

SELECT specimen_type_id,name
FROM `specimen_type` 
WHERE name LIKE 'whole%'

Any word that contains Urine

Get the ID of urine

SELECT specimen_type_id, name

FROM `specimen_type`

WHERE name LIKE '%Urine%'

Test Type

- Get all test types assigned to urine specimen type

SELECT t.test_type_id, t.name

FROM `test_type` t

INNER JOIN specimen_type s ON s.specimen_type_id = t.test_category_id

WHERE s.name LIKE 'urine'

- Get all test types assigned to Chemestry department

SELECT t.test_type_id, t.name
FROM `test_type` t
INNER JOIN test_category c ON c.test_category_id = t.test_category_id
WHERE c.name LIKE '%Chemestry%'

Measures

Any word that contains Potassium

Get ID for measure potassium

SELECT measure_id, name

FROM `measure`

WHERE name LIKE '%Potassium%'

Test Type Measure

- Get all measures for test type Electrolyte

SELECT tm.test_type_id, tm.measure_id, t.name AS testname, m.name AS measurename

FROM `test_type_measure` tm

INNER JOIN test_type t ON t.test_type_id = tm.test_type_id

INNER JOIN measure m ON m.measure_id = tm.measure_id

WHERE t.name LIKE 'Electrolyte'